|  | HYTEC ELECTRONICS Ltd HEAD OFFICE: 5 CRADOCK ROAD, READING, BERKS. RG2 0JT, UK Telephone: +44 (0) 118 9757770 Fax: +44 (0)118 9757566 E-mail: sales@hytec-electronics.co.uk |
| --- | --- |

IP-ADC-8413

16 Channel 16 Bit ADC

Product Specification

Document Nos: ADC 8413/PS/1.1

Date: 11/07/2006

CONTENTS

1.	Description	3

2.	Overall Specifications	3

3.	Operating Modes	4

4.	ADC Register Read Out	4

5.	Application Registers	4

5.1	Control & Status Register (CSR)	5

Control	5

5.2	Trigger Sample Number Register LS	5

5.3	Trigger Sample Number Register MS	6

5.4	Clock Rate	6

5.5	Vector	6

5.6	ADC Pre-trigger FIFO	6

5.7	ADC Post-trigger FIFO	7

5.8	FIFO Fullness Counter	7

5.9	ADC Registers	7

5.10	Auxilary Control Register (ACR)	8

Control	8

6.	ADC Operation	9

6.1	STROBE input as 10MHz clock signal and ADC sampling	9

6.2	FIFOs and Interrupts	9

6.2.1	Trigger Sample Numbers	9

6.2.2	Software Trigger	9

6.2.3	External Hardware Trigger and Clock	9

6.3	Register Update	9

7.	ID PROM	10

8.	Calibration	11

8.1	ID Page 1 (PG1 = 0, PG0 = 1) Calibration Values Layout.	12

8.2	ID Page 2 (PG1 = 1, PG0 = 0) Calibration Values Layout.	13

8.3	ID Page 3 (PG1 = 1, PG0 = 1) Calibration Values Layout.	14

9.	Isolation	15

10.	I/O Connector – 50 way on 8413 ADC Board	15

11.	Transition Card Connections	16

Connectors 1-4	16

# Description

The Hytec IP-ADC-8413 is an Industry Pack that provides 16 channels of simultaneously sampled analogue digitisation with the following characteristics:-

- 16 independent channels (one ADC per input)

- 16 bits resolution – 15 bits no missing codes

- Single full-scale trim for hardware gain adjustment.

- Software calibration by software driver possible using stored offset and gain parameters.

- True full differential inputs.

- +/-10V full-scale standard programmable to +/-5V full-scale resolution all inputs.

- Front-end instrumentation amplifiers can be factory set for gains of up to x1000.

- FIFO memories for single sample and triggered sample readout (256K post-trigger samples)

- Low offset error - +/- 2.5mV without software calibration. (+/-2LSBs after software calibration)

- Low gain error - +/- 0.5% FS without software calibration.(+/-2LSBs after software calibration)

- Low error drift  -  2ppm per deg C

- High input impedance – 1Gohms.

- Up to 160KHz sampling rate

- Simultaneous sampling – 70ns aperture delay time. Uncertainty time and channel matching 3ns.

- System to plant isolation to 100V when externally powered by DC/DC converter option

- Serial number, PCB issue and firmware issue held in ID PROM

- 8/32MHz system clock operation

- EPICS driver support

# Overall Specifications

Size:																Single width Industry Pack 1.8ins x 3.9 ins

Operating temp:							0 to 45 deg C ambient

Number of channels:			16

ADC resolution:						16 bits

Diff. Non-linearity:				Monotonic to 15 bits (at 160kHz throughput)

Int. Non-linearity:					+/-2LSBs max.

Offset error:										+/-2.5mV uncorrected.

Offset drift:										+/-0.5ppm per deg C typical

Gain error:											+/-0.5% uncorrected

Gain drift:											+/-2 ppm per deg C typical

Range:														+/-10V full-scale (+ve input referred to –ve input.

Cross-talk:											+/-1LSB channel to channel for FS input on adjacent channel.

CMRR														Greater than 80dB

CMV:														+/-12V.

Over-voltage:									+/-50V.

Throughput:										200KHz max

Aperture time:								70ns typical (conversion start to hold)

Conversion time:						3us (plus 1.6 us readout to register)

Bandwidth (-3dB):					100kHz (factory set – other cut-offs can be specified)

SNR:															-90dB at 1kHz typical

SINAD:													-90dB at 1kHz typical

Isolation:												100V via opto-isolators (if externally powered)

Data format:										16 bits two’s complement.

Memory:												Buffer register for each ADC conversion and FIFO for all 16 conversions

On-board FIFO:							256K conversion values with half full and full flags

Power:														+5V @ 300mA typical, +/-12V @ 200mA typical from VME or 8912.

# Operating Modes

There are three operating modes:-

- DC sampling – when the unit is armed or receives an External 10MHz signal, the inputs are sampled at the programmed clock rate and the conversions placed in a 16 conversion pre-trigger FIFO. Interrupt  is generated when the  FIFO is full. The FIFO may be readout as it is filling.

- Register mode – the last ADC reading may be read at random from each addressed ADC register.

- Triggered sampling  – When the board is triggered conversions are stored in a larger 256K conversion post-trigger FIFO. Interrupt Request is generated when the FIFO is full. The FIFO may be readout as it is filling. The pre-trigger FIFO remains unchanged. A record of the sample number when the trigger occurred is stored and can be read.

# ADC Register Read Out

There are sixteen ADC buffer registers (addresses 10hex – 2Ehex) which store the last sampled conversions and may be read at any time.  The channel order is channel 1 at address 10hex to channel 16 at address 2E.  Additionally there are two additional ADC registers to monitor the 0V Reference  (address 30Hex) and 2.5V Reference  (address 32Hex).

| D15 | D14 | D13 | D12 | D11 | D10 | D09 | D08 | D07 | D06 | D05 | D04 | D03 | D02 | D01 | D00 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| MSB |  |  |  |  |  |  |  |  |  |  |  |  |  |  | LSB |

| Range (RGE) bit from ACR Register (34hex) | 2’s Complement (2C) bit from ACR Register (34hex) | Range | ADC Value Negative Full Scale | ADC Value Positive Full Scale |
| --- | --- | --- | --- | --- |
| 0 | 0 | -10V to +10V | 8000h | 7FFFh |
| 0 | 1 | -10V to +10V | 0000h | FFFFh |
| 1 | 0 | -5V to +5V | 8000h | 7FFFh |
| 1 | 1 | -5V to +5V | 0000h | FFFFh |

# Application Registers

There are five application specific (I/O) registers; the CSR, the number of samples per trigger, the clock rate, the interrupt vector value and FIFO fullness register.  There are also 16 ADC buffer registers.

## Control & Status Register (CSR)

## Control

Write Address:  0 hex

| D15 | D14 | D13 | D12 | D11 | D10 | D09 | D08 | D07 | D06 | D05 | D04 | D03 | D02 | D01 | D00 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ARM | ET | ST | XC | EG | Ext | ETF | EF | EHF | QF | IC | RST | x | x | x | x |

Status

Read Address:  0 hex

| D15 | D14 | D13 | D12 | D11 | D10 | D09 | D08 | D07 | D06 | D05 | D04 | D03 | D02 | D01 | D00 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ARM | ET | ST | XC | EG | Ext | ETF | EF | EHF | QF | IC | RST | THF | FE | TF | F |

X = Not Used

ARM	When set, arm the ADCs. Allows sampling if EG bit is not set.

ET  	Enable trigger.  If set allows external trigger or software trigger to route sampled conversions to the post-trigger FIFO. The pre-trigger FIFO will retain the previously sampled conversions. The sample number will be stored in the trigger sample number register when a trigger occurs. The action is synchronised to the first sample clock after the rising edge of trigger.

ST  			Software trigger. Allows trigger action to be initiated by software command.

XC	External Clock

EG	Enable Go. When Strobe* is released sampling can commence according to the state of ET.

Ext  	If set to 0, the internal 10MHz clock is used to derive the sample rate. If set to 1 the external clock (Strobe*) is used.

ETF	Enables interrupt when the post-trigger FIFO is full. (Optional - normally set to 0)

EF	Enables interrupt when the pre-trigger FIFO is full.

EHF	Enables interrupt when the pre-trigger FIFO is half full.

QF	     The post-trigger FIFO is a quarter full. (Overflow of fullness counter.) Reset by THF +TF.

IC				When set to 1 (Internal Clock) generates 10MHz clock.

RST 	Resets the FIFO when set to 1.

THF	The post-trigger FIFO is half full.

FE				The pre-trigger FIFO is empty

TF	The post-trigger FIFO is full.

F	The pre-trigger FIFO is full. The FIFO contains the last ADC conversions. Cleared when FIFO is read or when a new conversion is read.

## Trigger Sample Number Register LS

Read/write Address:  2 hex

The least significant word of the sample number stored when trigger occurs

| D15 | D14 | D13 | D12 | D11 | D10 | D09 | D08 | D07 | D06 | D05 | D04 | D03 | D02 | D01 | D00 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| T15 | T14 | T13 | T12 | T11 | T10 | T9 | T8 | T7 | T6 | T5 | T4 | T3 | T2 | T1 | T0 |

## Trigger Sample Number Register MS

Read/write Address:  4 hex

The most significant word of the sample number stored when trigger occurs

| D15 | D14 | D13 | D12 | D11 | D10 | D09 | D08 | D07 | D06 | D05 | D04 | D03 | D02 | D01 | D00 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| T31 | T30 | T29 | T28 | T27 | T26 | T25 | T24 | T23 | T22 | T21 | T20 | T19 | T18 | T17 | T16 |

## Clock Rate

Read/write Address:  6 hex

The clock rate register is a four bit register which enables codes 0 – 16 to generate frequencies of 1 Hz to 200kHz in multiples of 1,2,5 or 10. (E.g. 0=1Hz, 1=2Hz, 2=5Hz, 3=10Hz and so on, see table below) Each clock pulse will initiate simultaneous ADC conversions and store them in memory.

Clock Rate Register Frequency Table

| RegisterValue | Clock Rate Frequency | RegisterValue | Clock Rate Frequency |
| --- | --- | --- | --- |
| 0 | 1Hz | 9 | 1KHz |
| 1 | 2Hz | 10 | 2KHz |
| 2 | 5Hz | 11 | 5KHz |
| 3 | 10Hz | 12 | 10KHz |
| 4 | 20Hz | 13 | 20KHz |
| 5 | 50Hz | 14 | 50KHz |
| 6 | 100Hz | 15 | 100KHz |
| 7 | 200Hz | 16 | 160KHz |
| 8 | 500Hz |  |  |

## Vector

Read/write Address:  8 hex

The vector register is a 16 bit register which stores the interrupt vector value.

| D15 | D14 | D13 | D12 | D11 | D10 | D09 | D08 | D07 | D06 | D05 | D04 | D03 | D02 | D01 | D00 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| V15 | V14 | V13 | V12 | V11 | V10 | V9 | V8 | V7 | V6 | V5 | V4 | V3 | V2 | V1 | V0 |

## ADC Pre-trigger FIFO

Read Address:  a hex

Read the pre-trigger FIFO memory (16 conversions). Reset Full and set FE when emptied..

| D15 | D14 | D13 | D12 | D11 | D10 | D09 | D08 | D07 | D06 | D05 | D04 | D03 | D02 | D01 | D00 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| A15 | A14 | A13 | A12 | A11 | A10 | A9 | A8 | A7 | A6 | A5 | A4 | A3 | A2 | A1 | A0 |

## ADC Post-trigger FIFO

Read Address:  c hex

Read the post-trigger FIFO memory (256K conversions). Reset TF and set TFE when emptied.

| D15 | D14 | D13 | D12 | D11 | D10 | D09 | D08 | D07 | D06 | D05 | D04 | D03 | D02 | D01 | D00 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| A15 | A14 | A13 | A12 | A11 | A10 | A9 | A8 | A7 | A6 | A5 | A4 | A3 | A2 | A1 | A0 |

## FIFO Fullness Counter

Read Address:  e hex

Two up/down counters which count conversions as they are entered / read from the pre-trigger FIFO (least significant byte ) and post-trigger FIFO (most significant byte). At the end of each trigger/readout sequence the value in the registers should be zero.  Since the capacity is 64K it should be used in conjunction with post trigger FIFO quarter full (QF), half full (THF) for 128K and full (TF) for 256K.

| D15 | D14 | D13 | D12 | D11 | D10 | D09 | D08 | D07 | D06 | D05 | D04 | D03 | D02 | D01 | D00 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| PO7 | PO6 | PO5 | PO4 | PO3 | PO2 | PO1 | PO0 | PR7 | PR6 | PR5 | PR4 | PR3 | PR2 | PR1 | PR0 |

## ADC Registers

Read/write Address:  10hex – 32hex

The first sixteen ADC buffer registers store the last sample conversions and may be read at any time.  The seventeenth is used to monitor the 0V reference and the eighteenth is used to monitor the 2.5V reference or in table form…

| Register Address (Hex) | Description | Register Address (Hex) | Description |
| --- | --- | --- | --- |
| 10 | ADC 0 | 22 | ADC 9 |
| 12 | ADC 1 | 24 | ADC 10 |
| 14 | ADC 2 | 26 | ADC 11 |
| 16 | ADC 3 | 28 | ADC 12 |
| 18 | ADC 4 | 2A | ADC13 |
| 1A | ADC 5 | 2C | ADC 14 |
| 1C | ADC 6 | 2E | ADC 15 |
| 1E | ADC 7 | 30 | 0V Reference |
| 20 | ADC 8 | 32 | 2.5V Reference |

| D15 | D14 | D13 | D12 | D11 | D10 | D09 | D08 | D07 | D06 | D05 | D04 | D03 | D02 | D01 | D00 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| A15 | A14 | A13 | A12 | A11 | A10 | A9 | A8 | A7 | A6 | A5 | A4 | A3 | A2 | A1 | A0 |

## Auxilary Control Register (ACR)

## Control

Read/write Address:  34hex

| D15 | D14 | D13 | D12 | D11 | D10 | D09 | D08 | D07 | D06 | D05 | D04 | D03 | D02 | D01 | D00 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| X | x | x | x | x | x | x | x | x | x | PG1 | PG0 | 2C | N/S | RST | RGE |

X = Not Used

RGE	Sets the range of the ADCs. 0 - +/- 10V and 1 - +/-5V.

RST			0 – Normal Operation, 1 – Reset ADC.

N/S			0 – Standby Mode, 1 – Normal Operation.

2C			  0 – ADC Values in 2’s Complement (default), 1 – ADC Values 0000 (Neg FS)-FFFF (Pos FS).

PG0	Bit 0 of ID PROM Paging.*

PG1	Bit 1 of ID PROM Paging.*

*Since some carrier devices only support 64 locations in the ID PROM, and the 8413 has up to 80 16-bit calibration values. Therefore it is required to page the ID PROM, these two bits are used to switch between pages of the ID PROM.

| PG1 | PG0 | PAGE | NOTES |
| --- | --- | --- | --- |
| 0 | 0 | 0 (Default) | Normal VITA4 format for ID PROM. |
| 0 | 1 | 1 | Calibration Values for ADC Channel 0-5. |
| 1 | 0 | 2 | Calibration Values for ADC Channel 6-11. |
| 1 | 1 | 3 | Calibration Values for ADC Channel 12-16. |

# ADC Operation

## STROBE input as 10MHz clock signal and ADC sampling

In systems which use multiple 8413s and 8004 carrier boards one of the 8004s is set up as a master card using the 8413 IP card in site A as the controlling card. This card is set up to generate 10MHz clock on DMAReq1 line which is routed to the 8004 front panel GO Lemo. This is connected to the INHIBIT front panel Lemo of the next 8004. The slave 8004 can be set so that INHIBIT is connected to its GO output thus transmitting the 10MHz on to the next 8004 card.

## FIFOs and Interrupts

One sample clock causes conversion of the 16 ADCs and the resulting conversion values fill the pre-trigger FIFO. Its output will indicate Full showing that a sample has occurred and the FIFO may be read and emptied. This may be indicated to the carrier board by setting IntReq0*. As the FIFO is being read the ADCs can convert a new sample. Thus throughput is limited by the block read time.

If the external FIFO option is fitted, in triggered mode, sampling and readout occur as for continuous mode until a trigger (external hardware or software generated) is received. Conversions are then written to the post-trigger FIFO which has space for 256K conversions. When this has been filled it will indicate Full and interrupt if enabled. Complete readout of either FIFO will restore it to an empty state as indicated by FIFO empty. The FIFOs may be reset at any time by writing 0 to their Full control bits.

### Trigger Sample Numbers

When a trigger occurs the current sample number is stored in a 32-bit register. The register can be reset to zero by writing to it. The sample number is reset whenever the Strobe* signal goes low.

### Software Trigger

The unit can be triggered by a software trigger by writing a ‘1’ to the Software Trigger (ST) bit of the CSR. The trigger state remains asserted (i.e. conversions are steered to the post-trigger FIFO) until the ST bit is cleared .

### External Hardware Trigger and Clock

The external trigger and clock are passed to the 8413 via designated pins (see Appendices B, C and D).

## Register Update

All ADC registers are updated simultaneously.

# ID PROM

As some IP Carrier Cards only support 64 16-Bit locations in the ID Space, we page the ID Space to provide extra space for ADC calibration data.   To switch between pages, there are control bits in the Auxilary Control Register (ACR). The default setting is the standard ‘VITA4’ layout, but there are additional pages as shown below…

| PG1 | PG0 | PAGE | NOTES |
| --- | --- | --- | --- |
| 0 | 0 | 0 (Default) | Normal VITA4 format for ID PROM. |
| 0 | 1 | 1 | Calibration Values for ADC Channel 0-5. |
| 1 | 0 | 2 | Calibration Values for ADC Channel 6-11. |
| 1 | 1 | 3 | Calibration Values for ADC Channel 12-16. |

The word addresses are as below:-

Base+80			ASCII	‘VI’												5649h

Base+82			ASCII	‘TA’												5441h

Base+84			ASCII	‘4 ‘													3420h

Base+86			Hytec ID high byte						0080h

Base+88			Hytec ID low word						0300h

Base+8A		Model number										8413h

Base+8C		Revision															0101h   This shows PCB Iss 1 Xilinx V1

Base+8E		Reserved														0000h

Base+90 		Driver ID														0000h

Base+92			Driver ID														0000h

Base+94			Flags																	0002h

Base+96 		No of bytes used								001Ah

Base+98 		Cal Type														xxxxh    0 = No Calibration, 1,2 = Calibration factors Stored.

Base+9A 		Serial Number										xxxxdec

Base+9C 		Not used														xxxxh

Base+9E 		WLO																	5555h

# Calibration

The type of calibration factors held in the ID PROM are specified by the Cal Type held at Base+98 in the ID PROM:-

0 = No Calibration factors held in ID PROM.

1 = 3 Point Calibration factors Stored in ID PROM.

2 = 5 Point Calibration factors Stored in ID PROM.

The Calibration Factors are held in the ID PROM pages 1 to 3 as shown the following Tables and as described in SECTION 7.  These are the stored ADC values, derived from reading the ADC at the following specified voltages…

| Value | Calibration Voltage |
| --- | --- |
| nFS | -10V |
| nHS | -5V |
| zero | 0V |
| pHS | +5V |
| pFS | +10V |

These values can then be used in the following equations to correct the offset and gain errors of the individual cannels of the ADC8413 IP card.

## ID Page 1 (PG1 = 0, PG0 = 1) Calibration Values Layout.

The layout of the calibration pages is additionally dependant on the ‘Cal Type’ Value from the Default (Page 0) page of the ID PROM Section.

| Address | Cal Type = 1 | Cal Type = 2 |
| --- | --- | --- |
| Base+80 | ADC0 Cal data nFS | ADC0 Cal data nFS |
| Base+82 | ADC0 Cal data Zero | ADC0 Cal data nHS |
| Base+84 | ADC0 Cal data pFS | ADC0 Cal data Zero |
| Base+86 | ADC1 Cal data nFS | ADC0 Cal data pHS |
| Base+88 | ADC1 Cal data Zero | ADC0 Cal data pFS |
| Base+8A | ADC1 Cal data pFS | ADC1 Cal data nFS |
| Base+8C | ADC2 Cal data nFS | ADC1 Cal data nHS |
| Base+8E | ADC2 Cal data Zero | ADC1 Cal data Zero |
| Base+90 | ADC2 Cal data pFS | ADC1 Cal data pHS |
| Base+92 | ADC3 Cal data nFS | ADC1 Cal data pFS |
| Base+94 | ADC3 Cal data Zero | ADC2 Cal data nFS |
| Base+96 | ADC3 Cal data pFS | ADC2 Cal data nHS |
| Base+98 | ADC4 Cal data nFS | ADC2 Cal data Zero |
| Base+9A | ADC4 Cal data Zero | ADC2 Cal data pHS |
| Base+9C | ADC4 Cal data pFS | ADC2 Cal data pFS |
| Base+9E | ADC5 Cal data nFS | ADC3 Cal data nFS |
| Base+A0 | ADC5 Cal data Zero | ADC3 Cal data nHS |
| Base+A2 | ADC5 Cal data pFS | ADC3 Cal data Zero |
| Base+A4 |  | ADC3 Cal data pHS |
| Base+A6 |  | ADC3 Cal data pFS |
| Base+A8 |  | ADC4 Cal data nFS |
| Base+AA |  | ADC4 Cal data nHS |
| Base+AC |  | ADC4 Cal data Zero |
| Base+AE |  | ADC4 Cal data pHS |
| Base+B0 |  | ADC4 Cal data pFS |
| Base+B2 |  | ADC5 Cal data nFS |
| Base+B4 |  | ADC5 Cal data nHS |
| Base+B6 |  | ADC5 Cal data Zero |
| Base+B8 |  | ADC5 Cal data pHS |
| Base+BA |  | ADC5 Cal data pFS |

Table Key -

Cal data	nFS - Negative Full Scale Calibration Factor.

Cal data	nHS - Negative Half Scale Calibration Factor.

Cal data	Zero	- Zero (0 Volts) Calibration Factor.

Cal data	pHS - Positive Half Scale Calibration Factor.

Cal data	pFS - Positive Full Scale Calibration Factor.

## ID Page 2 (PG1 = 1, PG0 = 0) Calibration Values Layout.

The layout of the calibration pages is additionally dependant on the ‘Cal Type’ Value from the Default (Page 0) page of the ID PROM Section.

| Address | Cal Type = 1 | Cal Type = 2 |
| --- | --- | --- |
| Base+80 | ADC6 Cal data nFS | ADC6 Cal data nFS |
| Base+82 | ADC6 Cal data Zero | ADC6 Cal data nHS |
| Base+84 | ADC6 Cal data pFS | ADC6 Cal data Zero |
| Base+86 | ADC7 Cal data nFS | ADC6 Cal data pHS |
| Base+88 | ADC7 Cal data Zero | ADC6 Cal data pFS |
| Base+8A | ADC7 Cal data pFS | ADC7 Cal data nFS |
| Base+8C | ADC8 Cal data nFS | ADC7 Cal data nHS |
| Base+8E | ADC8 Cal data Zero | ADC7 Cal data Zero |
| Base+90 | ADC8 Cal data pFS | ADC7 Cal data pHS |
| Base+92 | ADC9 Cal data nFS | ADC7 Cal data pFS |
| Base+94 | ADC9 Cal data Zero | ADC8 Cal data nFS |
| Base+96 | ADC9 Cal data pFS | ADC8 Cal data nHS |
| Base+98 | ADC10 Cal data nFS | ADC8 Cal data Zero |
| Base+9A | ADC10 Cal data Zero | ADC8 Cal data pHS |
| Base+9C | ADC10 Cal data pFS | ADC8 Cal data pFS |
| Base+9E | ADC11 Cal data nFS | ADC9 Cal data nFS |
| Base+A0 | ADC11 Cal data Zero | ADC9 Cal data nHS |
| Base+A2 | ADC11 Cal data pFS | ADC9 Cal data Zero |
| Base+A4 |  | ADC9 Cal data pHS |
| Base+A6 |  | ADC9 Cal data pFS |
| Base+A8 |  | ADC10 Cal data nFS |
| Base+AA |  | ADC10 Cal data nHS |
| Base+AC |  | ADC10 Cal data Zero |
| Base+AE |  | ADC10 Cal data pHS |
| Base+B0 |  | ADC10 Cal data pFS |
| Base+B2 |  | ADC11 Cal data nFS |
| Base+B4 |  | ADC11 Cal data nHS |
| Base+B6 |  | ADC11 Cal data Zero |
| Base+B8 |  | ADC11 Cal data pHS |
| Base+BA |  | ADC11 Cal data pFS |

Table Key -

Cal data	nFS - Negative Full Scale Calibration Factor.

Cal data	nHS - Negative Half Scale Calibration Factor.

Cal data	Zero	- Zero (0 Volts) Calibration Factor.

Cal data	pHS - Positive Half Scale Calibration Factor.

Cal data	pFS - Positive Full Scale Calibration Factor.

## ID Page 3 (PG1 = 1, PG0 = 1) Calibration Values Layout.

The layout of the calibration pages is additionally dependant on the ‘Cal Type’ Value from the Default (Page 0) page of the ID PROM Section.

| Address | Cal Type = 1 | Cal Type = 2 |
| --- | --- | --- |
| Base+80 | ADC12 Cal data nFS | ADC12 Cal data nFS |
| Base+82 | ADC12 Cal data Zero | ADC12 Cal data nHS |
| Base+84 | ADC12 Cal data pFS | ADC12 Cal data Zero |
| Base+86 | ADC13 Cal data nFS | ADC12 Cal data pHS |
| Base+88 | ADC13 Cal data Zero | ADC12 Cal data pFS |
| Base+8A | ADC13 Cal data pFS | ADC13 Cal data nFS |
| Base+8C | ADC14 Cal data nFS | ADC13 Cal data nHS |
| Base+8E | ADC14 Cal data Zero | ADC13 Cal data Zero |
| Base+90 | ADC14 Cal data pFS | ADC13 Cal data pHS |
| Base+92 | ADC15 Cal data nFS | ADC13 Cal data pFS |
| Base+94 | ADC15 Cal data Zero | ADC14 Cal data nFS |
| Base+96 | ADC15 Cal data pFS | ADC14 Cal data nHS |
| Base+98 |  | ADC14 Cal data Zero |
| Base+9A |  | ADC14 Cal data pHS |
| Base+9C |  | ADC14 Cal data pFS |
| Base+9E |  | ADC15 Cal data nFS |
| Base+A0 |  | ADC15 Cal data nHS |
| Base+A2 |  | ADC15 Cal data Zero |
| Base+A4 |  | ADC15 Cal data pHS |
| Base+A6 |  | ADC15 Cal data pFS |
| Base+A8 |  |  |
| Base+AA |  |  |
| Base+AC |  |  |
| Base+AE |  |  |
| Base+B0 |  |  |
| Base+B2 |  |  |
| Base+B4 |  |  |
| Base+B6 |  |  |
| Base+B8 |  |  |
| Base+BA |  |  |

Table Key -

Cal data	nFS - Negative Full Scale Calibration Factor.

Cal data	nHS - Negative Half Scale Calibration Factor.

Cal data	Zero	- Zero (0 Volts) Calibration Factor.

Cal data	pHS - Positive Half Scale Calibration Factor.

Cal data	pFS - Positive Full Scale Calibration Factor.

# Isolation

The ADC 8413 +/-12 volt power supply can be derived either internally (non-isolated) from the carrier card (VME +/-12V) or from external isolating DC-DC converters (type 8912) mounted on an 8211 transition card.  The source is selected using jumpers J1, J2 and the GND-AGND link.

J1		External +12V connect 1 & 2,  Internal +12V	connect 2 & 3

J2		External -12V  connect 1 & 2,  Internal -12V	connect 2 & 3

J3 	Reserved

J4		Factory set (boot jumper)

GND-AGND Link

IN for internal +/-12V (non-isolated)

OUT for external +/-12V  (isolated and supplied from transition card DC-DC converter).

# I/O Connector – 50 way on 8413 ADC Board

| Pin | Signal | Pin | Signal |
| --- | --- | --- | --- |
| 1 | Input 1 + | 26 | Input 13 - |
| 2 | Input 1 - | 27 | Input 14 + |
| 3 | Input 2 + | 28 | Input 14 - |
| 4 | Input 2 - | 29 | Input 15 + |
| 5 | Input 3 + | 30 | Input 15 - |
| 6 | Input 3 - | 31 | Input 16 + |
| 7 | Input 4 + | 32 | Input 16 - |
| 8 | Input 4 - | 33 |  |
| 9 | Input 5 + | 34 |  |
| 10 | Input 5 - | 35 | XTrig+ |
| 11 | Input 6 + | 36 | XTrig- |
| 12 | Input 6 - | 37 | N.C. |
| 13 | Input 7 + | 38 | N.C. |
| 14 | Input 7 - | 39 | XClk+ |
| 15 | Input 8 + | 40 | XClk- |
| 16 | Input 8 - | 41 | +12VX |
| 17 | Input 9 + | 42 | AGND |
| 18 | Input 9 - | 43 | +12VX |
| 19 | Input 10 + | 44 | AGND |
| 20 | Input 10 - | 45 | -12VX |
| 21 | Input 11 + | 46 | AGND |
| 22 | Input 11 - | 47 | -12VX |
| 23 | Input 12 + | 48 | AGND |
| 24 | Input 12 - | 49 | AGND |
| 25 | Input 13 + | 50 | AGND |

# Transition Card Connections

TB 8213 I/O Connector – 50 way on transition

# Connectors 1-4

| Pin | Signal | Pin | Signal |
| --- | --- | --- | --- |
| 1 | Inp1 - | 26 | Inp1 + |
| 2 | Inp2 - | 27 | Inp2 + |
| 3 | Inp3 - | 28 | Inp3 + |
| 4 | Inp4 - | 29 | Inp4 + |
| 5 | Inp5 - | 30 | Inp5 + |
| 6 | Inp6 - | 31 | Inp6 + |
| 7 | Inp7 - | 32 | Inp7 + |
| 8 | Inp8 - | 33 | Inp8 + |
| 9 | Inp9 - | 34 | Inp9 + |
| 10 | Inp10 - | 35 | Inp10 + |
| 11 | Inp11 - | 36 | Inp11 + |
| 12 | Inp12 - | 37 | Inp12 + |
| 13 | Inp13 - | 38 | Inp13 + |
| 14 | Inp14 - | 39 | Inp14 + |
| 15 | Inp15 - | 40 | Inp15 + |
| 16 | Inp16 - | 41 | Inp16 + |
| 17 |  | 42 |  |
| 18 | Xtrig - | 43 | Xtrig + |
| 19 |  | 44 |  |
| 20 | XClk - | 45 | XClk + |
| 21 | AGnd | 46 | +12V |
| 22 | AGnd | 47 | +12V |
| 23 | AGnd | 48 | -12V |
| 24 | AGnd | 49 | -12V |
| 25 | AGnd | 50 |  |