# Bug fixed for LeCroy scope application

- Reset scope when IOC boot up: we take out PINI from reset record;

- IOC reboot: in device support somehow memcpy was called with length equal –1, fixed;

- Auto calibration cause link lost: we can either disable ACAL or recover link on fly;

- Position bug when analyze read back: fixed;

- Each channel has own handler: RHIC’s way, we didn’t see any benefit;

- Scope doesn’t return whole response: reported by RHIC, we never saw it, but our driver can handle it;

- Waveform shows very strangely: in log mode, we changed;

- Waveform only return a quarter length: in device support memcpy need sizeof(float);

- Waveform doesn’t update anymore after we disable then re-enable channel: in device support, we forget to turn pact to false if read waveform failed;

- Reset scope when IOC boot up: we take out PINI from load panel record;

- WAVEDESC structure alignment issue: in driver, fixed;

- waveform return 0 length: in device support  don’t reset nord;

- IOC crash: in device support use WAIT_FOREVER is msgQreceive;

- Recover link failed: in driver we turn LINK_OK too late, fixed;

- mbuf issue;

- IOC crash: in device we have a hidden dead loop cause message receiver will send himself message;

- memPartInfoGet issue;

- IOC crash: in device support, process last_trgtime record twice on one event;